Cif File Entries for the Purdue Quest Diffractometer (June 2024)

A CIF (Crystallographic Information File) is written by the refinement program, but it is usually incomplete, and some required items need to be added by hand before the CIF file will pass checking. Checking, done online via the web page https://checkcif.iucr.org/, is required if you want to publish crystallographic data in a peer reviewed journal.

This document contains general data collection items for all data sets collect on the two Purdue Bruker Quest diffractometer.

The CIF format uses text only (ASCII characters). This means, if it is not on the computer keyboard, don’t use it.

Each basic entry consists of two items: The identifier, and the value. The identifier usually starts with an underscore, and might consist of several words connected by underscores (e.g. _diffrn_radiation_monochromator). It is followed by the value, which, depending on the item, is a number or text. It has to be either a single number or word, or it has to be marked as a single item with single parentheses (e.g. 'fine focus sealed tube'). Each line should not exceed 80 characters (that’s legacy from times when e-mails could not handle longer lines). If an entry is longer than 80 characters, you should use semicolons instead of the single parentheses.  It may then stretch over several lines.

Values have to be in the cif “library”. Numbers have to be in an acceptable range, words have to be one of a list of allowed entries in the “CIF Dictionary” (e.g. at http://www.iucr.org/resources/cif/dictionaries/cif_core).

Tables in a cif file are following a more complicated procedure.

For a complete description of the CIF file format for reporting crystal structures, and on current developments regarding cif files, see http://www.iucr.org/resources/cif.

The CIF code allows to express some non-ASCII characters (such as Greek letters or Umlauts). H\"ubschle, for example, will be printed as Hübschle when converted to text using “printcif” or making tables.

Beware of typos!!

With a CIF file based on a dataset from one of the Purdue Bruker Quest instruments and refined with Shelxle and Shelxl2019, usually the items listed on the next pages have to be added or edited. They are listed in the order in which they will appear in the cif file and are given in blocks similar as in the cif file. Items in black can be copied directly. Items in red need a second look and a value looked up or a decision made.

Edits can be made using a simple text editor such as Notepad or programs like Encifer or FinalCif. Do not use advanced text editors such as MS Word.

You may also try out FinalCif to automatically extract missing items from the raw data files and to create tables and reports (read the FinalCif software documentation first!).

1. Cell measurement details

_cell_measurement_reflns_used     from p4p file, 1st value of SAINGL line

_cell_measurement_theta_min       from p4p file, 2nd value of SAINGL line

_cell_measurement_theta_max       from p4p file, 3rd value of SAINGL line

2. Crystal colour and shape

_exptl_crystal_description        from p4p file, MORPH entry

_exptl_crystal_colour             from p4p file, CCOLOR entry

_exptl_crystal_size_max           from p4p file, CSIZE entry

_exptl_crystal_size_mid           from p4p file, CSIZE entry

_exptl_crystal_size_min           from p4p file, CSIZE entry

3. Absorption correction details

_exptl_absorpt_correction_type    'multi-scan'

_exptl_absorpt_correction_T_min   obtain value from “abs” file, 3rd last line

_exptl_absorpt_correction_T_max   obtain value from “abs” file, 3rd last line

_exptl_absorpt_process_details

;

SADABS 2016/2: Krause, L., Herbst-Irmer, R., Sheldrick G.M. & Stalke D. (2015). J. Appl. Cryst. 48, 3-10.

;

(These four entries might be already added if the original hkl file has been used rather than the one processed in XPREP).

4. Instrument details

Quest instrument with molybdenum source:

_diffrn_source                    'fine focus sealed tube X-ray source'

_diffrn_source_current            '30 mA'

_diffrn_source_voltage            '50 kV'

_diffrn_radiation_monochromator   'Triumph curved graphite crystal'

_diffrn_measurement_device        'three circle diffractometer'

_diffrn_measurement_device_type   'Bruker AXS D8 Quest'

_diffrn_detector_type             'PhotonII charge-integrating pixel array (CPAD)'

_diffrn_measurement_method        'omega and phi scans'

_diffrn_detector_area_resol_mean  7.4074

Quest instrument with copper source:

_diffrn_source                    'I-mu-S 3.0 microsource X-ray tube'

_diffrn_source_current            '1.1 mA'

_diffrn_source_voltage            '50 kV'

_diffrn_radiation_monochromator   'HELIOS multilayer Montel optics'

_diffrn_measurement_device        'four circle diffractometer'

_diffrn_measurement_device_type   'Bruker AXS D8 Quest'

_diffrn_detector_type             'PhotonIII_C14 charge-integrating and photon counting pixel array'

_diffrn_measurement_method        'omega and phi scans'

_diffrn_detector_area_resol_mean  7.4074

The entries _diffrn_radiation_monochromator, _diffrn_measurement_device, _diffrn_detector_type, _diffrn_source_current and _diffrn_source_voltage are not automatically added to the cif file. Add them anyway (it also helps to populate all entries in reports you might create using FinalCIF).

5. Software details:

_computing_data_collection        'Apex5 v2023.9-2 (Bruker, 2023)'

_computing_cell_refinement        'SAINT V8.40B (Bruker, 2020)'

_computing_data_reduction         'SAINT V8.40B (Bruker, 2020)'

_computing_structure_solution     'SHELXS-97 (Sheldrick, 2008)' or 'SHELXM (Sheldrick, 2008)' or 'SHELXT (Sheldrick, 2015b)'

_computing_structure_refinement

;

SHELXL-2019/2 (Sheldrick, 2015a, 2019),

SHELXLE Rev1573 (H\"ubschle et al., 2011)

;

6. Structure solution methods

_atom_sites_solution_primary      'direct', 'patt', 'dual space' or 'isomor'

_atom_sites_solution_secondary    'difmap' or 'none'

Use 'direct' if you used TREF in XS or XM to solve the structure (direct methods). Use 'patt' if you used PATT in XS to solve the structure. Use 'dual space' if you used ShelXT to solve the structure. Use 'isomor' if you solved the structure by isomorphous replacement.

Use difmap if you assigned Q-peaks as atoms. Use none if no atoms needed to be added during refinement.

7. Absolute Configuration:

_chemical_absolute_configuration 'ad' or 'syn' or 'unk'

This entry only applies for chiral and non-centrosymmetric structures.

Use 'ad' (anomalous dispersion) if you refined a Flack parameter and got a reliable value close to zero with a small esd.

Use 'syn' (synthesis) if you assigned the absolute structure based on a known chiral center, e.g. from the starting material that did not change during synthesis.

Use 'unk' (unkown) if you don’t know.

8. Special Details:

If you used any non-standard techniques to process the data or refine the structure, add these to the _refine_special_details section of the cif. Add here details on twinning, disorder, any constraints or restraints you used, pseudosymmetry, use of Squeeze and other details you would like to add that fit nowhere else.

_refine_special_details

;

Typical example text:

The two disordered moieties were restrained to have similar geometries.

Uij components of ADPs for disordered atoms closer to each other than

2.0 Angstrom were restrained to be similar. Subject to these conditions

the occupancy ratio refined to ??? to ???.

Water H atom positions were refined and O-H and H...H distances were

restrained to 0.84(2) and 1.36(2) Angstrom, respectively. Some water

H atom positions were further restrained based on hydrogen bonding

considerations.

Water H atom positions were initially refined and O-H and H...H

distances were restrained to 0.84(2) and 1.36(2) Angstrom,

respectively, while a damping factor was applied. Some water

H atom positions were further restrained based on hydrogen bonding

considerations. In the final refinement cycles the H atoms were

constrained to ride on their carrying oxygen atom (AFIX 3) and

the damping factor was removed.

(see the cif template files for other example text)

;

9. Authors, software cited above and acknowledgements. Add below the data_ line at the top and above the remainder of the cif text. Repeat the author loop (below _publ_author_address) as often as needed to include all authors (you, advisor, etc). Delete any software you did not use.

_publ_contact_author_name       'Last Name, First Name(s)'

_publ_contact_author_address

;       Your Department

Purdue University

Your Street Address

West Lafayette

IN Your ZIP-code

USA

;

_publ_contact_author_email      'your e-mail'

_publ_contact_author_fax        'your department fax number'

_publ_contact_author_phone      'your department phone number'

loop_

_publ_author_name

_publ_author_address

'Last Name, First Name(s)'

;       Department

University or Company

Street Address

City

State, ZIP-code

Country

;

_publ_section_references

;

Bruker (2023). Apex5 v2023.9-2, SAINT V8.40B, Bruker AXS

Inc.: Madison (WI), USA, 2023.

H\"ubschle, C. B., Sheldrick, G. M. and Dittrich, B. (2011). J. Appl. Cryst.

44, 1281--1284.

Krause, L., Herbst-Irmer, R., Sheldrick G.M. & Stalke D. (2015). J. Appl.

Cryst. 48, 3--10.

Parsons, S, Flack, H., Wagner, T. (2013). Acta Cryst. B69, 249-259.

Sheldrick, G.M. (2008). Acta Cryst. A64, 112--122.

Sheldrick, G. M. (2008). CELL_NOW. Version 2008/4. University of G\"ottingen,

Germany.

Sheldrick, G. M. (2019). SHELXL2019. University of G\"ottingen, Germany.

Sheldrick, G. M. (2015a). Crystal structure refinement with SHELXL, Acta Cryst.

C71, 3--8.

Sheldrick, G. M. (2015b). SHELXT--Integrated space-group and

crystal-structure determination, Acta Cryst. A71 3--8.

Sheldrick, G. M. (2012). TWINABS. Ver. 2012/1.

Spek, A. L. (2003). J. Appl. Cryst. 36, 7--13.

van der Sluis, P. & Spek, A. L. (1990). Acta Cryst. A46, 194-201.

;

_publ_section_acknowledgements

;

This material is based upon work supported by the National Science Foundation

through the Major Research Instrumentation Program under Grant No. CHE 1625543.

(Funding for the single crystal X-ray diffractometer).

We thank the College of Science, Purdue University, for support for X-ray

source and detector upgrades through the 2020 and 2023 Laboratory and

University Core Facility Research Equipment Program.

;